## Slide 1
Green/Black 1
Fisher to Breakout box cable wire assignment
Black/Green 2
Black/Brown 3
Brown/Black 4
Green/Red 5
Red/Green 6
Black/Orange 7
Orange/Black 8
Black/Yellow 9
Yellow/Black 10
White/Black 11
Black/White 12
White/Red 13
Red/White 14
Red/Black 15
Black/Red 16
Red/Orange 17
Orange/Red 18
Red/Brown 19
Brown/Red 20
Black/Blue 21
Blue/Black 22
Red/Yellow 23
Yellow/Red 24
Blue/Red
Red/Blue
Unused

## Slide 2
Breakout Box
Top of Fridge
MC plate
Glenair
Fischer
BNC plugs
Blue: custom attocube cables
Purple: custom D-SUB to Fischer cables
BNC cable
DVM
(USED TO MEASURE RESISTANCE VALUE)
DVM
(USED TO MEASURE RESISTANCE VALUE)
How resistance values are measured

## Slide 3
DC Line Resistances from Glenair on MC plate to either breakout box BNC plugs (Loom 5 and 6) or end of attocube cables (special loom) (Hence, the resistance values include the wires made)
Loom 5 | R (ohms) | Loom 6 | R (ohms) | Loom Special | R (ohms)
1 | 22 | 25 | 22 | 1 | 4
2 | 22 | 26 | 21 | 2 | 4
3 | 22 | 27 | 22 | 3 | 4
4 | 21 | 28 | 21 | 4 | 4
5 | 22 | 29 | 22 | 5 | 4
6 | 21 | 30 | 21 | 6 | 4
7 | 22 | 31 | 22 | 7 | NC
8 | 21 | 32 | 22 | 8 | 13
9 | 22 | 33 | 22 | 9 | 14
10 | 21 | 34 | 21 | 10 | 13
11 | 22 | 35 | 22 | 11 | 14
12 | 21 | 36 | 21 | 12 | 13
13 | 22 | 37 | 22 | 13 | 14
14 | 21 | 38 | 21 | 14 | 13
15 | 22 | 39 | 22 | 15 | 14
16 | 21 | 40 | 21 | 16 | 13
17 | 22 | 41 | 22 | 17 | 14
18 | 20 | 42 | 21 | 18 | 13
19 | 22 | 43 | 22 | 19 | 14
20 | 21 | 44 | 21 | 20 | 13
21 | 22 | 45 | 22 | 21 | 14
22 | 20 | 46 | 21 | 22 | 13
23 | 22 | 47 | 22 |  | 
24 | 21 | 48 | 21 |  | 
4
3
2
1
5
6
7
8
Special low resistance loom
Fisher 5 and 6 used.

## Slide 4
5.84mm
1.49mm
0.33mm
0.58mm
0.58mm
500nm diameter (not drawn to scale)
5um
The sharper the tip, the better
We need the side of the tuning fork and the side of the tip to be flat. (embedded in the same plane)